Name:											  Date:

Competency Statement: The sterile processing team member has completed facility or health care organization-required education and competency verification activities related to performing steam sterilization.1

- Guideline for sterilization. In: Guidelines for Perioperative Practice. Denver, CO: AORN, Inc.

Outcome Statement: Patient is free from signs and symptoms of infection.2

- Petersen C, ed. Infection. In: Perioperative Nursing Data Set. 3rd ed. Denver, CO: AORN, Inc; 2011:254-276.

| Competency Statements/Performance Criteria | Verification Method [Select applicable code from legend at bottom of page] | Not Met (Explain why) |
| --- | --- | --- |
|  | DEM/ DO/DA | KAT | S/SBT/ CS | V | RWM/ P&P | O |  |
| Selects saturated steam under pressure for heat and moisture stable items unless otherwise indicated by the manufacturer. |  |  |  |  |  |  |  |
| Follows manufacturer’s written instructions for use (IFU) for operating steam sterilizers. |  |  |  |  |  |  |  |
| 3. Determines correct cycle parameters by reconciling cycle parameters recommended by the device manufacturer with the sterilizer manufacturer’s written IFU for the specific sterilization, or |  |  |  |  |  |  |  |
| following the device manufacturer’s written IFU when the sterilizer and device manufacturer’s written IFU cannot be reconciled. |  |  |  |  |  |  |  |
| Verifies that load configuration and placement of items within the sterilizer is in compliance with the sterilizer manufacturer’s written IFU. |  |  |  |  |  |  |  |
| 5. Uses physical monitors (eg, printouts, graphs) for every load and reviews to verify that conditions necessary for steam sterilization have been met. |  |  |  |  |  |  |  |
| 6. Uses external and internal chemical indicators with each package to verify that parameters for sterilization have been achieved according to the manufacturer’s written IFU. |  |  |  |  |  |  |  |
| Places a class I chemical indicator (ie, process indicator) on the outside of every package unless the internal indicator is visible through the package material. |  |  |  |  |  |  |  |
| Places a class 5 chemical indicator (ie, integrating indicator) or a class 6 chemical indicator (ie, emulating indicator) inside every package. |  |  |  |  |  |  |  |
| Places a class 3 or class 4 chemical indicator within a package to meet the requirements for internal monitoring when indicated. |  |  |  |  |  |  |  |
| Places chemical indicators in an area within the package that represents a challenge for air removal and steam contact. |  |  |  |  |  |  |  |
| Consults the chemical indicator manufacturer, the device manufacturer, and the container manufacturer when there is a question concerning the correct number or placement of internal chemical indicators. |  |  |  |  |  |  |  |
| Uses biological indicators at least weekly to monitor sterilizer efficacy and for load release purposes according to the manufacturer’s written IFU. |  |  |  |  |  |  |  |
| Does not release loads containing an implant until the result of the biological indicator test is available. |  |  |  |  |  |  |  |
| Removes items from the steam sterilizer and leaves them untouched until cool enough to handle without concern that retained moisture may act as a wick for bacteria on the hands of personnel who touch the package. |  |  |  |  |  |  |  |
| Allows items to cool on the sterilization rack. |  |  |  |  |  |  |  |
| Does not place warm or hot items on cool or cold surfaces. |  |  |  |  |  |  |  |
| Recognizes that sterile packages that have formed condensate are considered unsterile and the contents should not be used. |  |  |  |  |  |  |  |
| Documents steam sterilization cycles and modalities accurately, completely, and legibly according to facility or health care organization policies and procedures including a. load identification (eg, patient name, items processed, operator information); |  |  |  |  |  |  |  |
| b. type of cycle, date and time of cycle, and cycle parameters; |  |  |  |  |  |  |  |
| c. contents; and |  |  |  |  |  |  |  |
| d. results of physical, chemical, and biological monitors. |  |  |  |  |  |  |  |
| Verbalizes a review of facility or health care organization policies and procedures related to steam sterilization. |  |  |  |  |  |  |  |
| Participates in quality improvement programs related to steam sterilization as assigned. |  |  |  |  |  |  |  |